GLOSS METER OPERATING

PROCEDURE

This gloss meter complies with ASTMD 523.85 and ASTM E97.

Identification of Terms

Gloss meter:	An instrument used to measure the amount of light reflected in a mirror like fashion.  60( gloss meter is common for all-purpose measurements.

Shiny Appearance:	This is when a beam of light is reflected from a surface is basically one direction (i.e.:  mirror on highly polished surface).

Flat or Matte Appearance:	This is when a beam of light is reflected from a surface in all directions (the light is diffused/scattered).

Gloss:	Amount of light reflected in a mirror like fashion relative to the amount of light that would be reflected, under identical circumstances, by a highly polished plane surface of black glass having a refractive index of 1.567 (this refractive index is equivalent to a gloss value of 100).

Calibration:	Adjusting the gloss value reading of the gloss meter to a known standard (see gloss defined).

Procedure

- Inspect the under side (bottom surface) of the instrument to be sure that it is clean and free of scratches and rough particles.  A rough surface will damage the test piece and/or calibration standard.

- Inspect the lenses and their channels removing any dust or stray material.  A small camel’s hair brush should be used to clean the lenses and surrounding area.

- Verify calibration.

  - Wipe the calibration piece (polished black glass) clean using a soft lint free cloth.  DO NOT SCRATCH OR MAR THE SURFACE!  Refer to Calibration Standard Cleaning Care for more details.

  - Place the gloss meter on the standard.

  - Press and hold the Read button until the display reading stabilizes.

  - Verify the reading.

    - Standard is 93.2 ( 0.3.  Acceptable range 92.9 - 93.5 gloss units.

    - With the read button depressed adjust the calibration to read within the required parameters.

  - Take 2nd reading to verify.

Note:  Once the gloss meter has been calibrated be careful not to adjust the calibrating dial as this will affect the calibration.

- Perform required gloss measurements:

  - Place the gloss meter on the test specimen.  Specimen must be flat and smooth in order for proper readings to take place.

  - Periodically confirm the calibration of the meter during the testing desired.

Maintenance

- Battery charging

  - Battery is to be recharged only after the “LO BAT” symbol is displayed.  This will ensure maximum battery pack life.

  - The unit will be charged in approximately 16 hours.

- Lamp changing

Refer to the manual for details and specific instructions.

- Calibration Standard cleaning & care.

  - Use only dry lint - free cloth to clean the standard.

  - Do not use detergents or any cleaning agents that may leave a residue.

  - Do not scratch or mar the surface of the standard!!