MICROMETER

## CALIBRATION PROCEDURE #2

ENVIRONMENT:

The environment must be relatively dust free.  Clean surrounding inspection area with Citri-Clean, clean all inspection tools.

## A.	VISUAL INSPECTION

- Inspect the casting body, ratchet, thimble and the sleeve for evidence of damage.

- Inspect the anvil and spindle for wear, burrs.

- Inspect the protective carrying case for cleanliness, cracks and wear.

- Using WD-40 clean the micrometer (disassemble if extremely dirty or seized).  Do not oil the carbide contact points.

- Re-assemble.

- Fully close the micrometer.

- Inspect the function of the micrometer.

- Does the anvil and spindle align properly?

- Does the spindle rotate smoothly?

- Does the spindle lock clamp properly?  (Adjust or repair if necessary).

- Does the ratchet stop or friction thimble work properly?

- Are the numbers and lines clearly readable?

## B.	CALIBRATION

- Cleaning the storing oil from the gage blocks.

- Clean the gauge block stand.

- Measure the OD of .250, .500, .750, 1.000 gage blocks.  1 –2 “micrometers and larger, require the gauge blocks to be wrung together.  Do this by applying a small amount of molycoat lubricant to one measuring surface, rub the surfaces of each gauge block together.  Apply pressure, the two surfaces will adhere.

- Record the actual measurements on the Inspection Report.